## Slide 1: Principles of Radiation Detection
Spectrometeric detectors can be divided into two groups
  Those using ionizing phenomenon of the radiation: Pulse Ionization Chambers, Proportional Counters, Semiconductor Counters/Detectors.
  Those using scintillation process: Scintillation Detectors.
Ionization Based Detectors
  The quantum or particle interacts with the sensitive volume of the detector material and produces a specific number of ion pairs N which is proportional to the energy E of the quantum/particle.
  With the help of electrodes, an electric field is established.
  Ions are collected at the electrodes charging the interelectrode capacitance.
  The charge leads to a voltage V = Q/C = Nq/C = qE/eC where N is number of ions, q is the charge of an ion, e is the energy needed to create an electron-ion pair, and E is the energy of the particle.

## Slide 2: Pulsed Ionization Chamber
A pulsed ionization chamber is a hermetically sealed vessel filled with gas in which two electrodes produce a strong electric field.
The negative electrode ( as a substrate) is coated with a thin layer of radioactive compound.
Electrodes may be planar, spherical or cylindrical.
Grid with a small positive potential helps in screening the electrons to the positive electrode. When the electron moves between the grid and collecting (positive) electrode, the collected charge is qN since they pass through identical potential difference.
+ Collecting Electrode
- Electrode
Grid
d

Particle or Gamma
Radioactive material

## Slide 3: Collimated Detection
+ Anode
- Cathode
Quantum
Collimator
+
-
ions

## Slide 4: Ionizing Chambers
Common gases filled such as N2 or Ar2 have oxygen impurities which creates errors because electrons “stick” to the heavy molecules of the electronegative gas and are slowed down to their motion to the collecting electrodes. This is due to fact that the mobility of such ions is lower than that of electrons.
  This effect can be minimized by having a mixture of nitrogen (~96%) and argon (~4%) gases. In this case, electron energies correspond to the minimum capture cross-section and the sticking (to the negative ion) effect becomes small.
Other sources of errors include ion recombination and scatter of particles within the chamber.
  It is necessary to use sufficiently large electric field to reduce recombination errors.
  Pulse chambers are usually employed to measure the energies of heavy charged particles such as alpha particles.

## Slide 5: Proportional Counters
By increasing the voltage on the electrode of an ionization chamber, electrons in the electrical field can be accelerated to affect the ionization, thus increasing the number of electrons reaching at the collecting electrode. This creates a multiplication factor M0.
Impact ionization of gas molecules is accompanied by their excitation and subsequent de-excitation and emission of ultraviolet photons.These ultraviolet photons create a photoelectric effect at the cathode to knock out secondary photoelectrons.
The secondary electrons overwhelm the photoelectric effect.
They are usually of cylindrical form because it permits to have a high potential gradient near the collecting electrode as it is smaller in dimensions (at center).
In the presence of impurities, linearity may be affected.
Gas multiplication coefficient can be obtained from unity to 10,000.
They are more efficient than simple pulse ionization chamber

## Slide 6: Semiconductor Detectors
The particle energy is transformed into electric pulses at the junction region of semiconductor material: silicon, germanium,..
In p-n junction silicone detectors, the useful region or the ionization chamber is located on the very surface of the detector. the p-type thin layer ( of donor material such as phosphorus) is constructed over the n-type silicon crystal.
Holes in p region diffuse through the n-regions, and electron in n-regions diffuse in p-regions. Thus, p-region is negatively charged and n-region is positively charged creating an electric potential barrier to further diffusion of holes and electrons. This depletion layer can be increased by applying positive voltage to the n-region and negative voltage to the positive region to create a high potential gradient in the depletion layer.
If an electron-hole pair is created by interaction of a quanta in the depletion layer, electron will move to wards positive electrode via n-silicon layer and holes towards negative electrode via p-region.
Thus a voltage is produced through the charging the capacitor V = qN/C. The thickness of depletion layer is analogous to the diameter of ionization chamber.

## Slide 7: Advantages of Semiconductor Detectors
Much smaller e (minimum energy needed to create an electron-hole pair) ~3.6 eV for silicon. This leads to better detector resolution.
Independence of e is also realized in terms of mass and charge pf the particle. This due to the strong electric potential gradient across the depletion layer.
Small charge collection time (<10 ns).
Very small recombination losses due to fast charge collection.
PROBLEMS: For correct measurement, the thickness of the depletion layer must be much larger than the range of measured particles.
For high energy particles (alpha particles), the range is workable. For example, at 5MeV, the range is 20 m This leads to a detector made out of silicon with resistivity of 500 ohm-cm at a base voltage of 100V.
For a particle of energy 100 keV, range is ~150 m. a very thick depletion layer is required which is difficult. For energies above 100 keV in X-rays or gamma rays, Germanium-Lithium Ge(Li) is used.

## Slide 8: Scintillation Detectors
Consists of a scintillation phosphor and a photomultiplier coupled together through an optical contact.
Fast charged particles or gamma rays interact with the scintillation material to excite molecules which returns to their ground stae emitting optical photons.. The intensity of each scintillation event (amount of light) is proportional to the energy lost by the particle or gamma ray in the phosphor.
Photomultiplier tubes are used to amplify the optical photon intensity and produce voltage proportional to the energy of the particle creating scintillation.
Photons arriving at the photo-sensitive photocathode of the PM tube produce photoelectrons knocking out the photoelectron from the cathode.The number of electrons emitted by the photocathode N0= kg0wE/Eph where g0 is the coefficient to account for losses at the photocathode, w is quantum efficiency of PM, Ephis the average energy of photon.
Finally, primary emission stimulates the secondary emission due to high potential gradient, working as a positive feedback providing a multiplication factor. N= kg0wmE/Eph where N is prepositional to the energy.

## Slide 9: Scintillation Detector
Advantages
  High detector efficiency
  Large amplitude and output signal
  High Speed
Most commonly used materials: alkali-halide crystals
  NaI(T1)
  CsI(T1)
  BGO (Bismuth Germinate), and BaF
Organic crystals: anthracene and stilbene
Conversion efficient is almost independent of energy in the range of .01 to 10 MeV. For NaI(T1),  (minimum energy) is 0.65 eV.
Efficiency is /square root of E in MeV. where  = 8.8 square root of .

## Slide 10: Scintillating Materials
Material density index of decay const light yield
(g/cm3) refraction (ns) (quanta/keV)
NaI(Tl) 3.67 1.85 230 38
BGO 7.13 2.15 300 8.2
LSO 7.4 1.82 40 15-27
BFC 408 1.032 1.59 5 12

## Slide 11: Pulse Shapes of the detectors
Let us consider a detector circuit equivalence as
R
C

## Slide 12: Pulse Shapes of Semiconductor Detectors
With respect to N= number of electron-hole pairs, and mobility of holes p, the maximum field Wmax, the voltage across the junction U, we can show

## Slide 13: Pulse Shape of Scintillating Detectors
A particle passing through the phosphor of a scintillation counter, produces excited molecule which emit optical photons during the decay or turning back to ground state. The total number of excited molecules n0 = kE/Eph. The decay for emission of optical photon is similar to radioactive decay with a time constant t0. The number of photons are then multiplied by PM tube. If N0 is the number of photons reaching at the anode, the decay of photon emission is given by